HS-GC-MS

The sample solutions to analyze were transferred to 20 mL HS glass vials and crimp-capped with a PTFE/Silicone cap (MACHEREY-NAGEL GmbH & Co. KG, Düren, Germany). Samples were measured on a GC-MS system from Thermo Fisher Scientific GmbH (Bremen, Germany) consisting of a TRACE™ 2000 GC with an S/SL injector installed and a Polaris Q ion trap mass spectrometer. The scan range was set between 30 and 400 m/z. The transfer line temperature was set to 300 °C, the ion source temperature to 200 °C and damping gas flow was set to 0.3 mL/ min helium. 3 micro scans were recorded at a maximum ion time of 25 ms. After initial 2 min at 40 °C, the GC oven temperature was raised to 100 °C with 10 °C/min and then to 200 °C with 50 °C/min. The injector was operated in split mode at 200 °C with a flow rate of 10 mL/min, the split ratio was 10 and the column flow 1 mL/min. As column, a Zebron ZB-1MS 60 m from Phenomenex (Aschaffenburg, Germany) was used, with 0.25 mm inner diameter and 1 µm film thickness. Prior to injection, the sample vials were incubated at 40 °C for 30 s while shaking in a CombiPAL autosampler (CTC Analytics AG, Zwingen, Switzerland).

For Headspace:

The syringe temperature was set to 50 °C and analysis was started by injecting 500 µL of the headspace content.

For SPME

Prior to measurement the grey SPME fiber (Divinylbenzene/Carboxen/Polydimethylsiloxane) 30 µm film thickness (supelco, merck, Darmstadt Germany) was conditioned at 200°C for one hour. For each measurement the SPME fiber was exposed to the sample vial for 10 min and the desorbed in the injector for additional 10 min.